SDK - HTMarch.dll Manual

VB 6.0 IDE

Note:

HTMarch.dll was compiled under VC++6.0.

The following ifdef block is the standard way of creating macros which make exporting

from a DLL simpler. All files within this DLL are compiled with the HTMARCH_API

symbol defined on the command line. this symbol should not be defined on any project

that uses this DLL. This way any other project whose source files include this file see

HTMARCH_API functions as being imported from a DLL, wheras this DLL sees symbols

defined with this macro as being exported.

#ifndef  HTMARCH_API

#define  HTMARCH_API  extern "C" __declspec(dllimport)

#endif

#define WIN_API __stdcall

Function Introduction

- Function declaration:

HTMARCH_API short WIN_API  dsoOpenDevice(unsigned short DeviceIndex)

Return value: Return zero (0) indicates device isn’t connected; return one (1) indicates device connected.

Parameter:

DeviceIndex

The first connected device index is 0, and others sequentially numbered.

Remark:

The device whose device index value is judged as DeviceIndex whether connected to PC or not.

Program  example:

unsigned short nDev = 0;

if(dsoOpenDevice(0) == 1)

{

;// Device connected

}

Else

{

;// Not detect device

}

- Function declaration:

HTMARCH_API unsigned short WIN_API dsoChooseDevice(unsigned short DeviceIndex, short nType);

Return value: Return zero (0) indicates failure; return one (1) indicates success.

Parameter

DeviceIndex

indicates current device index value.

nType

0:  logic analyzer Hantek6022BL

1:  Hantek6022BE

Remark:

Choose device

3. Function declaration:

HTMARCH_API short WIN_API dsoSetVoltDIV(unsigned short DeviceIndex,int nCH,int nVoltDIV);

Return value: one (0) for setup success and zero (0) for failure.

Parameter:

DeviceIndex

indicates current device index value.

nCH:

Channel index value. 0 stands for CH1, and 1 stands for CH2.

nVoltDIV：

Voltage index value. Minimum voltage is 0, and following is the index values for corresponding voltages.

0: 20mV/DIV

1: 50mV/DIV

2: 100mV/DIV

3: 200mV/DIV

4: 500mV/DIV

5: 1V/DIV

6: 2V/DIV

7: 5V/DIV

Remark:

The device whose device index value is judged as DeviceIndex whether connected to PC or not.

Programme example:

dsoSetVoltDIV(0,0,5);// Set the voltage of CH1 to 1V/DIV.

4. Function declaration:

HTMARCH_API short WIN_API dsoSetTimeDIV(unsigned short DeviceIndex,int nTimeDIV);

Return value: one (0) for setup success and zero (0) for failure.

Parameter

nDeviceIndex

indicates current device index value.

nTimeDIV

indicates current sampling rate index value, following is the value.

0 ~ 10 : 48MSa/s

11: 16MSa/s

12: 8MSa/s

13: 4MSa/s

14 ~ 24: 1MSa/s

25: 500KSa/s

26: 200KSa/s

27: 100KSa/s

Remark:

Setup device sampling rate.

Programme example:

5. Function declaration:

HTMARCH_API short WIN_API dsoReadHardData(

unsigned short DeviceIndex,

short* pCH1Data,

short* pCH2Data,

unsigned long nReadLen,

short* pCalLevel,

int nCH1VoltDIV,

int nCH2VoltDIV,

short nTrigSweep,

short nTrigSrc,

short nTrigLevel,

short nSlope,

int nTimeDIV,

short nHTrigPos,

unsigned long nDisLen,

unsigned long * nTrigPoint,

short nInsertMode);

Return value: Reading data, return “-1” for failure and non “-1” for success.

Parameter:

unsigned short DeviceIndex: Device index value

short* pCH1Data: CH1 data storage buffer pointer

short* pCH2Data: CH2 data storage buffer pointer

unsigned long nReadLen: The length of reading data

short* pCalLevel: Proofreading level (reference function dsoGetCalLevel)

int nCH1VoltDIV: The voltage of CH1

int nCH2VoltDIV: The voltage of CH2

short nTrigSweep: SWP MODE-0: AUTO; 1: Normal; 2: Single

short nTrigSrc: Trigger source - 0: CH1; 1: CH2

short nTrigLevel: Trigger level - 0 ~ 255

short nSlope: Trigger Slope - 0: Rise; 1: Fall

int nTimeDIV: Sampling rate

short nHTrigPos: Horizontal trigger position -0 ~ 100

unsigned long nDisLen: The length of the display data

unsigned long * nTrigPoint: The index value of returned trigger point

short nInsertMode: D-value mode - 0: Step D-value; 1: Line D-value; 2: SinX/X D-value

Remark:

Call this function to read data.

6. Function declaration:

HTMARCH_API unsigned short WIN_API dsoGetCalLevel(unsigned short DeviceIndex,short* level,short nLen);

Return value: return zero (0) for success and non-zero for failure.

Parameter:

DeviceIndex

indicates current device index value.

level

Proofreading data storage buffer.

nLen

The length of proofreading data, here=32.

Remark:

Acquire proofreading data from device.

Programme example:

short nCal[32];

dsoGetCalLevel(0, nCal l,32);

7. Function declaration:

HTMARCH_API short WIN_API dsoCalibrate(unsigned short nDeviceIndex,int nTimeDIV,int nCH1VoltDIV,int nCH2VoltDIV,short* pCalLevel);

Return value: return zero (0) for success and non-zero for failure.

Parameter:

nDeviceIndex

indicates current device index value.

nTimeDIV

Sampling rate

nCH1VoltDIV

The voltage of CH1

nCH2VoltDIV

The voltage of CH2

pCalLevel

Proofreading data memory area

Remark:

When any channel’s zero reference offset, it’s able to call this function to calibrate, and the calibration information stored at pCalLevel. If no offset, no need to call this function.

8. Function declaration:

HTMARCH_API unsigned short dsoSetCalLevel(unsigned short DeviceIndex,short* level,short nLen);

Return value: return zero (0) for success and non-zero for failure.

Parameter:

DeviceIndex

indicates current device index value.

level

Proofreading data memory area

nLen

The length of proofreading data, here is 32.

Remark:

After calling the function-dsoCalibrate to calibrate, it’s able to call this function to store the acquired calibration data to device for future directly reading purpose.

Programme example:

short nLevel[32];

dsoCalibrate(0,11,5,5,nLevel); // When zero reference offset, calibrate it first, then acquire the calibration data.

dsoSetCalLevel(0, nLevel,32); // Store the calibration date to device.